## Slide 1
Cal/Val of GEMS Level 1 Data:
Current Status and Inflight Approaches
Yeeun Lee1,2), Myoung-Hwan Ahn1,2), Mina Kang3,4), Junghyun Seo5), Hyunkee Hong5), Jaehoon Jeong5), Wonjin Lee5), Dai Ho Ko6), Jhoon Kim7)
1) Severe Storm Research Center (SSRC), Ewha Womans Unversity, Republic of Korea
2) Department of Climate and Energy Systems Engineering, Ewha Womans Unversity, Republic of Korea
3) City College of New York, USA
4) NOAA-Cooperative Science Center for Earth System Sciences and Remote Sensing Technologies, USA
5) Environmental Satellite Center, National Institute of Environmental Research (NIER), Republic of Korea
6) Korea Aerospace Research Institute (KARI), Republic of Korea
7) Department of Atmospheric Science, Yonsei University, Republic of Korea
August 19, 2025
[TEMPO-GeoXO Workshop, SAO 2025]

## Slide 2: GEMS: MORE THAN FIVE YEARS IN OPERATION
Solar Irradiance
(WSD & RSD)
Dark Current & LED
z
Earth Radiance
GEMS data monitoring system
  Since its launch in February 2020:
Solar diffuser
Scan mirror
Telescope
Spectrometer
Detector
Electronics
…
Radiation
(radiance)
+) Temperature information
: Input data for Level 2 retrievals
Issues related to:
1) Solar irradiance only (inversely to reflectance)
2) Both solar irradiance & Earth radiance
DC

## Slide 3: GEMS: MORE THAN FIVE YEARS IN OPERATION
Version history of Level 1 data (solar irradiance)
Year | Operation | Input Level 1 | Description | Improvement | Issues
2020 | V1.2.4 ( V2.0 in the NIER sftp) |  | Operational data |  | BTDF Systematic bias
2023 | V1.3.0 | EOSRL WSD | Empirical correction for the BTDF and simple degradation correction | BTDF correction Degradation correction | Intensified stripping features Residuals in N-S dependence
2024 Early | V1.4.0 | EOSRL WSD | Adjust the spectral sampling interval and resolution of the correction factor + Systematic correction (+10%) | De-striping | Residuals in N-S dependence & stripping Degradation
2024 Late | V1.5.0 | NIER WSD | Updated BTDF correction + bias correction (+10%) | Minimizing N-S dependence < 1% Reducing striping artifacts | TBU (Now evaluating the impact on L2)
2025 | TBU | NIER WSD & RSD | BTDF & diffuser degradation correction + new bias correction (BTDF definition) | De-trending | TBU
Year | Operation | Input Level 1 | Description | Improvement | Issues
2020 |  |  | Operational data |  | BTDF Systematic bias
2023 | V1.3.0 | EOSRL WSD | Empirical correction for the BTDF and simple degradation correction | BTDF correction Degradation correction | Intensified stripping features Residuals in N-S dependence
2024 Early | V1.4.0 | EOSRL WSD | Adjust the spectral sampling interval and resolution of the correction factor + Systematic correction (+10%) | De-striping | Residuals in N-S dependence & stripping Degradation
2024 Late | V1.5.0 | NIER WSD | Updated BTDF correction + bias correction (+10%) | Minimizing N-S dependence < 1% Reducing striping artifacts | TBU (Now evaluating the impact on L2)
2025 | TBU | NIER WSD & RSD | BTDF & diffuser degradation correction + new bias correction (BTDF definition) | De-trending | TBU

## Slide 4: DEGRADATION: DIFFUSER + OPTICS/ELECTRONICS
The primary diffuser, WSD (daily observation) The secondary diffuser, RSD (every 6 months
(exposure times = ~1800) (exposure times = 13)
Residual degradation
Solar irradiance changes @ 320 nm
  GEMS has the two solar diffuser, the primary (WSD, left), and the secondary diffuser (RSD, right).
  Solar diffusers are degraded most significantly compared to other optics.
* Solar azimuth angle

## Slide 5: DEGRADATION: DIFFUSER + OPTICS/ELECTRONICS
Solar irradiance changes @ 320 nm
Deviation of WSD/RSD from the initial observations [%]
5
In WSD/RSD,
common degradation perfectly cancel out.
-> Only diffuser degradation remain
* Solar azimuth angle

## Slide 6: DEGRADATION: DIFFUSER + OPTICS/ELECTRONICS
% differences along the N-S direction at each wavelength for WSD & RSD
Sep 21, 2020, WSD Sep 22 2020, RSD Sep 23 2024, WSD Sep 22, 2024, RSD
4 yrs passed
Solar irradiance changes
  Along spectral & spatial (N-S) directions
* Solar azimuth angle
Solar Irradiance
Dark Current & LED
z
Earth Radiance
Solar diffuser
Scan mirror
Telescope
Spectrometer
Detector
Electronics
…
(radiance)

## Slide 7: DEGRADATION: DIFFUSER + OPTICS/ELECTRONICS
Diffuser degradation correction:
Ratio to
the initial observation
W: The primary solar diffuser (daily)
R: The secondary solar diffuser (bi-annually)
(Ludewig et al., 2020)
: Methodology established -> Optimization in progress
Diffuser degradation correction

## Slide 8: SYSTEMATIC BIAS – TSIS COMPARISON
Spectral difference of GEMS solar irradiances compared to TSIS
Sep 15, 2020 () Sep 16, 2024 ()
Increase by ~15% ()
Bias in solar irradiances
  Radiometric calibration
  BTDF definition (refer to the presentation of Dr. Heesung Chong, SAO)
4 yrs passed

## Slide 9: SYSTEMATIC BIAS – VALIDATION APPROACHES
GEMS (Hyperspectral)
GEO
LEO
AMI
GOCI-2
OMPS
TROPOMI
MODIS
Higher availability for ray-matching
At the same 1) time
With same 2) spatial and 3) spectral response
With same 4) viewing geometry
View an
identical target
Collocation process
Inter-calibration approach
  GEO (twin satellites – GK2A & GK2B) & hyperspectral instrument onboard LEO
Potential for spectral analysis

## Slide 10: SYSTEMATIC BIAS – VALIDATION APPROACHES
Evaluation with LEO sensors (TROPOMI & OMPS)
  2021 –2025, every five days / Deep convective clouds (DCC)
Sensor | Target (# /month) |  |  | VZA | V(S)ZA |  | 
TROPOMI (=1~3%) | DCC (~4000) | < 15 min. | < 1.8 km (The half of LEO FOV) | < 1% | < 40 | < 3% | > 0.75
OMPS (=2%) | DCC (~100) |  | < 25 km (The half of LEO FOV) |  | < 40 |  | > 0.75
Sensor | Target (# /month) |  |  |  | V(S)ZA |  | 
 | DCC (~4000) | < 15 min. | < 1.8 km (The half of LEO FOV) | < 1% |  | < 3% | > 0.75
 | DCC (~100) |  | < 25 km (The half of LEO FOV) |  |  |  | > 0.75
Collocation thresholds
The number of collocation samples

## Slide 11: SYSTEMATIC BIAS – VALIDATION APPROACHES
Evaluation with LEO sensors (TROPOMI & OMPS)
Solar irradiance
Earth radiance
In 2021

## Slide 12: SYSTEMATIC BIAS – VALIDATION APPROACHES
Evaluation with LEO sensors (TROPOMI & OMPS)
Solar irradiance
Earth radiance
In 2024

## Slide 13: SYSTEMATIC BIAS – VALIDATION APPROACHES
Evaluation with LEO & GEO sensors (AMI/GK-2A – atmospheric imager)
Bias of GEMS compared to TROPOMI & AMI (all, DCCs)
@ 470 nm
Bias within 5% (@470 nm)
when compared to LEO &GEO Sensors

## Slide 14: SUMMARY & FUTURE WORKS
REMAINING ISSUES
  Solar irradiance only (inversely to reflectance):
    Diffuser BTDF N-S dependence : & Sinusodial pattern
    Diffuser degradation : -dependent
    Diffuser BTDF definition :
  Both solar irradiance & Earth radiance:
    Temporal discontinuity & striping :
    Throughput change : -dependent
    Radiometric calibration : -dependent
Degradation
Bias
Temporal variation
Degradation
Temporal variation
Bias

## Slide 15
Yeeun Lee: yeeunlee@ewha.ac.kr

## Slide 16: GEMS: five years of operation
Level 0-1B Process
Level 0
Co-addition Correction
Pixel Flagging
Crosstalk/
Nyquist Phase Correction
Offset/Linearity Correction
Smear Correction
Gain/
FPE Temperature Correction
Integration Time/PRNU
Dark Correction
Geometric Calibration
Stray Light Correction
Spectral Shift Correction
(Ir)radiance Calibration
2-Pixel Binning
L1B-Earth
Spatial Registration
Diffuser BTDF
Spectral Shift Correction
Spatial Registration
2-Pixel Binning
L1B-Solar
Green Box: updates
BTDF Correction
Degradation Correction
Empirical correction
(On-going update)

## Slide 17: GEMS Level 1C calibration updates:
Issue (prominent) | Irradiance (IRR) | Radiance (RAD) | Reflectance (REF) | Effect | Status
North-South dependence | O | X | O (inverse) | Systematic N-S bias | Reduced (<1%, REF)
Diffuser degradation | O | X | O (inverse) | Higher degradation @ shorter WVs | To be updated this year
Sensor degradation (post diffuser) | O | O | X | Degradation | -
BTDF Negative bias | O | X | O | Systematic negative bias | Removed (Courtesy: SAO)
Stripping | O | O | X | N-S spatial discontinuity | -
Reboot discontinuity | O | Potential | Potential | Temporal discontinuity | -
Radiometric calibration coefficient | O | O | X | Systematic bias | To be investigated this year
Straylight | X | O | O | Spectral dependence @ shorter WVs | -
* Updated only for IRR

## Slide 18: GEMS Level 1C calibration updates:
Solar irradiance (w/o BTDF update, V2.0)
Spatial anomaly [%]
Temporal variations at each spatial index [%]
@ 354 nm
@ 480 nm
* BTDF: Bi-directional Transmittance Distribution Function

## Slide 19: GEMS Level 1C calibration updates:
Solar irradiance (w/ BTDF update, V1.5.0)
Spatial anomaly [%]
Temporal variations at each spatial index [%]
Degradation (10% @ 354 nm )
Over 17% @ 320 nm
@ 354 nm
@ 480 nm

## Slide 20: GEMS Level 1C calibration updates:
GEMS diffuser degradation
z
Scan mirror
Solar diffuser
(WSD or RSD)
Solar Irradiance
Earth Radiance
Telescope
Spectrometer
Detector
Electronics
…
WSD:
RSD :
: Total degradation
: Diffuser contribution
: Degradation from common parts
: Temporal variations
Ludewig et al. (AMT 2020) & Kleipool et al. (AMT 2022)
Similar approach
for OMI and TROPOMI

## Slide 21: GEMS Level 1C calibration updates:
Stripe-like artifacts along north-south direction on the CCD
  Clearly appear in weak signals about 2% variations (i.e., dark ocean)
Radiance
March 20, 2022
* Potentially by non-linearity correction in the ADC process
Reflectance =

## Slide 22: GEMS Level 1C calibration updates:
Instrument reboot
  Occurs when GEMS is rebooted after entering survival mode, potentially related to electronic chain
@ 480 nm (Nadir)
% depending on WV
5/9, 2021

## Slide 23: GEMS Level 1C calibration updates:
Spatial anomaly of LED signals along the N-S direction (2024/09/22)
Optics degradation - LED observations:
Spatial anomaly [%]
LED Signals

## Slide 24: GEMS Level 1C calibration updates:
* After the spectral cal.
Spatial anomaly along the N-S direction (Sep 22, 2024)
DCs (co-addition) Electrons (b/f dark cal.) Electrons (a/f dark cal. ) a/f Radiometric cal.
Optics degradation – Level 0-1 calibration process