## Slide 1
Tropospheric Emissions:
Monitoring of Pollution
TEMPO Solar Calibration Status
Heesung Chong1, X. Liu1, J. Houck1, D.E. Flittner2, J. Carr3, W. Hou1, J. E. Davis1, R. M. Suleiman1, K. Chance1, J. Park1, G. González Abad1, C.R. Nowlan1, H. Wang1, J. Fitzmaurice1, B. Baker4, J. Lasnik4, D. Nicks4, L. Flynn5, S. Marchenko67, J. Joiner6, Colin Seftor67, Glen Jaross6
1SAO, CfA 2NASA LaRC 3Carr Astro. 4BAE 5NOAA/NESDIS 6NASA GSFC 7STC
June 15, 2026

## Slide 2: Solar dark time series (quadrant median)
1st safe mode 2023-09-08)
2nd safe mode (2024-01-24)
Change integration time (2024-10-10)
(1)
(2)
Linear trend since 2024/10/10
Quadrant A: 1.60%/yr
Quadrant B: 0.73%/yr
Quadrant C: 1.72%/yr
Quadrant D: 1.35%/yr
(3)
median

## Slide 3: Spectral Calibration Wavelength Shifts V3 vs. V4
Spatial index = 1025
A wavelength shift occurred between 07/31/2025 and 08/07/2025 (0.03–0.04 nm).
Two events during this period:
Idle: No radiance data collected between 07/31 and 08/06
Thermal control for FPA radiators was adjusted: 08/02

## Slide 4: Irradiance Variations
TEMPO solar irradiance variability = Diffuser degradation + Optics/detector degradation + Seasonal variability (goniometry errors) + Solar activity
The seasonal variability differs by wavelength.
Plan to determine each term: - Neglect the solar activity - Perform diffuser degradation correction first - Fit ‘optics/detector degradation’ and ‘seasonal variability’ simultaneously
Plan for V05 online solar irradiance processing: - Using pre-determined look-up tables, extrapolate (predict) the three terms and provide degradation- and seasonal-variability-corrected solar irradiance

## Slide 5: TEMPO solar irradiance changes
UV (290–490 nm)
Working irradiance change
Working/reference irradiance ratio change
(1-day before and after each reference measurement)
Working and reference irradiance changes
Both working and reference irradiance changes exhibit seasonal variations, and they do not fully cancel out in the ratio.
 To improve the fit, we eventually excluded the solstice measurements, which have the most extreme range of angles.

## Slide 6: The model, suggested by GEMS Team
Ratio
Exponential degradation of diffuser transmittance:
Degradation factor
(unit: 1/hr)
Asymptote
(dimensionless)
* Boundary condition:
at .
* Asymptote:
at .
21 spectral points averaged (~4 nm)
21 spectral points averaged (~4 nm)

## Slide 7: Characterizing reference diffuser
Working diffuser
Reference diffuser
Azimuth-angle dependence
Irradiance change
[%/degree]
So far, we have derived the working diffuser’s in-flight sensitivity only and applied it to the reference diffuser.
Now, the reference diffuser has its own look-up table.
Irradiance change [%]

## Slide 8: Updating reference diffuser table
Before
After
Elevation-angle dependence has not been updated because we have very few samples at different elevation angles.

## Slide 9: Checking spectral continuity
r
c
Reasonably good continuity from UV to VIS CCDs
40 nm spectral smoothing + some spatial averaging was used here

## Slide 10: Diffuser degradation correction results
Before correction
After correction
Difference
(optics/detector degradation) + (seasonal variability)
21 spectral points averaged (~4 nm) (Note that a 40-nm binning window was used to derive the degradation parameters.)
320.0 nm

## Slide 11: Summary
Dark current of CCDs has been better than projected pre-launch
  Presently is increasing at the modest rate of slightly less than 2%/yr
Wavelength scale has been stable since June 2024, except for…
  After returning to operating mode, a shift back towards First light by ~0.03 nm occurred.
Diffuser transmittance correction (BTDF)
  Artifacts in solar irradiance seasonal variability reduced between V3 and V4, but can do better
  Degradation of diffuser and optics/detector are evident, along with uncorrected goniometry
  Working on approach for V5 to correct for temporal changes involving:
    Diffuser degradation
    Optics/detector changes
    Uncorrected goniometry
  Will examine polarization correction for irradiance (IRR)
Impact of etalon structure in calibration + ground to orbit image shift reduced in V4, but the team will work to better represent the in-flight case

## Slide 12
Thank you for your attention!